CAUTION

SOP Autoclave room 105 APS

GET TRAINING IF YOU HAVE NOT USED THIS AUTOCLAVE BEFORE!

HINTS TO AVOID COSTLY REPAIRS:

- The machine is purposely set so that a minimum of buttons need to be pushed to start it. If you are turning knobs - DON'T. Random twiddling does not improve performance.

- DO NOT set the dry time to anything greater than 20 min. That is more than sufficient. More time is resulting in extreme wear (and repairs) on the S5 solenoid.

- There are 2 power switches - OCCASIONALLY they are turned off when repairs are being made or the building power is off - so if you arrive to no display, check the power is on first.

- ALWAYS put a tray under your load - and ALWAYS clean up any spills!

- If there is a problem (malfunction, water leaking, whatever) seek IMMEDIATE assistance - don't just walk away - PLEASE!

- The autoclave is under service contract. The manual is beside the autoclave, plus information on how to reach the service provider.

See below for:

Operation of the Autoclave

Hints and Troubleshooting

OPERATION OF THE AUTOCLAVE

BEFORE STARTING THE AUTOCLAVE LOAD:

- Is the door properly secured? - door is double hinged (apply pressure evenly to BOTH sides of door wheel when closing chamber door  CLOCKWISE)

- all the door “pins” have to lock up inside the outer part of the chamber door

- 4 cycles available - at Sterilize Temp 121C, cycles 1-3 are preset

cycle 1 - Liquid 	- Sterilize Time 25 min

cycle 2 - Gravity	- Sterilize Time 30 min, Dry Time 20 min

cycle 3 - Gravity	- (garbage: Biological samples) Sterilize Time 2 hr, Dry Time 20 min

cycle 4 - this cycle can be changed – follow directions below

IMPORTANT for cycle 4: change ONLY the type of cycle (LIQUID, OR GRAVITY and the TIME - NOTHING ELSE)

- press cycle number “once” - will give information on that cycle

- press cycle number “twice” within 5 sec - will start cycle

- Check that steam is not escaping around the door - it makes quite a loud noise so it should be obvious if there is a problem

- “reset” will abort cycle

- at end of cycle - open door SLOWLY about 2.5 cm, to allow steam to escape - autoclave begins 10 min countdown - MUST wait 10 minutes - even if door will open - need to get liquids below the boiling point of 100C by 5C - to prevent liquids exploding.

- IMPORTANT - let printer finish printing message - THEN hit reset (if too early hitting “reset” printer will just keep trying to print warning message)

- “auto utility” days in work week 5, shutdown 5 pm, restart  8:30 am

(if “auto utility on” - press any numbered button to activate the unit)

- during a scheduled steam shut down or campus closure (ie Christmas) - adjust power switch to “standby” - shuts off all steam and water

- SIGN the LOG BOOK - near the autoclave

BEFORE LEAVING THE AUTOCLAVE - YOUR LOAD IS COMPLETE

- Remove autoclaved material using gloves - return gloves to room 105

TEMPERATURE/STEAM SETTINGS:

3 Valves above the chamber door – HI/Lo, Steam and Water

DO NOT adjust the HI-LO steam setting.  Correct position - turns clockwise - doesn’t need to be completely open    Steam (red) and Water (blue) – Correct position - both should be all the way counter clockwise to the stop point

The AMSCO service technician has set the autoclave for LO steam (121C/250F)

The steam pressure available in the University lines is not sufficient to operate this autoclave at 132C/270F.

AUTOCLAVING GARBAGE:

- Please autoclave garbage at end of day - after 4:30 pm

- Use cycle 3 (gravity - sterilize time 120 min (2hr))

- Since our temperature is 121C, longer exposure times required for a complete kill.

- If possible - combine garbage loads with other users (do not stack bags).

- Always put a SHALLOW drip tray under the garbage bag.

- Do not seal the bags - steam MUST be able to get into and out of the bag.

- Remove completed loads promptly

- Seal bags for disposal.

- Autoclaved garbage smells badly - please consider the people who work in this lab.

- Autoclaving Sharps containers – room 105 has a gravity displacement unit – therefore the lid should be open and the container on its side – put the lid on after sterilization is complete.

AUTOCLAVE LOADING HINTS:

- Center load in autoclave (front to back and side to side) - Away from walls

- This is a STEAM STERILIZER - Sealed containers will NOT be sterilized

- Steam MUST be able to get into and out of each item in load

- Use shallow trays - higher sides on the trays mean longer exposure times required to get steam into load.

- NOTE - for loads that contain both dry & liquid materials, use the liquid cycle.

- Liquids

- DO NOT CAP containers tightly - pressure build up may cause glass may break and liquid also will not sterilize properly.

- vessel should be twice as large as volume to be autoclaved.

- DO NOT MOVE hot liquids - bumping may cause bottles to break explosively.

- Choose appropriate cycle/sterilize time for your load - keep in mind a longer exposure time is better than too short - refer to the binder for more information.

TROUBLE SHOOTING HINTS

Check strainer not plugged (ie if pressure not decreasing very fast, error “low temp”, chamber full of water -  this may be the problem)

- located front of chamber, center bottom

- threads in and out - NOTE - it will be HOT - wear gloves

- 2 slots on each side - need to put something solid across slots (screwdriver) to get it started

Control panel selections include - sterilize, exhaust, complete, door unlocked

If complete or door unlocked does not appear - consider the following:

- if only the lamps are burned out - the printer should cycle (make a whirling sound) when the exhaust cycle is complete (ie load completed) and the door opened - if no sound as the door is opened then the door sensor is stuck

- another way to check the door sensor - with the door open - where the radial arms push up and lock - look at the top of the chamber where the door would fit - will see a plate - push up on this plate with a screwdriver and release - should hear the printer cycle - if not then a door problem

- if the door sensor is not working properly - SAFETY CONCERN - reason - someone could accidently push cycle (ie gravity or liquid) - easily done as drying oven sits close and walkway not wide at that point - even with the door open, steam would be introduced into the chamber - ie 30 lb of steam rushing into the chamber - it would be extremely difficult to get to the reset switch to stop it without getting scalded.

Re cleaning the chamber

- do not attempt to clean a hot chamber – caution burn hazard

Door won’t open

If door won’t open counter clockwise - turn door handle to right a partial turn and then try counter clockwise again

Demand Print

By pressing “Print” button during a cycle, unit will print: D - time of day - temp and pressure - If more problems develop push the rocker button forward where it says “Print” to get more detailed information.